| Number: | FMEA-002 | Revision: | 3 |
| --- | --- | --- | --- |
| Title: | Failure Mode Effects Analysis |
| UNCONTROLLED DOCUMENT UNLESS ISSUED WITH A RED STAMP This document is the confidential property of AGM, Inc. and may not be reproduced without prior written consent. |
| Revision Approvals: | Date |
| Originator | Heather Gray | 1/16/04 |
| [Originating dept. approval] |  |  |
| [Other dept. approval] | Emily Miner | 1/16/04 |
| Quality Assurance | Erdrin Azemi | 1/16/04 |
| General Management | AGM, Inc. | 1/16/04 |

# Purpose

  - To determine the components of a device that could fail and to determine the risk assessment of that failure.

# Scope

  - The re-design of the filter will be based upon the efficiency of the filter components and their risk assessment.

# References

  - N/A

# Revision History

  - Revision 1:  Proper document formatting

  - Revision 2:  Update of document

| Function or Component | Failure Mode | Effect on System | Possible hazards | Risk Index | User Detection Means | Applicable Controls |
| --- | --- | --- | --- | --- | --- | --- |
| Delivery and Retrieval Sheath | User confusion between two sheaths | Inability to efficiently deploy or retrieve filter | Spasm of artery, risk of patient complication | 11 | Unable to properly deploy or retrieve filter | Proper labeling of package |
| Nitinol Ring | Inconsistent diameter size | Complete failure of filter | Thromboembolism to the brain, decrease of blood flow | 8 | Unabe to properly fit around the artery | Ensure user training, nitinol flexibility |
| Porous Netting | Pore size too large or too small | Decreases efficiency of filter | Thromboembolism to brain, no blood flow | 8 | Constriction of blood flow | More testing to ensure proper pore size |
| Filter | Lack of conformation to vessel | Decreases efficiency | Stuck in the stent, spasm if artery, Thromboembolism to the brain | 6 | Spasm of artery, decrease of blood flow | Use a more conformable material |
| Nitinol Ring | Lack of conformation to vessel | Decreases efficiency of filter | Causes thromboembolism to the brain | 6 | Loss of blood flow | Use a more conformable material able to expand under physiological conditions |