Universal Facemask Adapter

Failure Mode Effect Analysis (FMEA)

April 15, 2007

Nicole Burns

Mike Callahan

LaDahvia Flournoy

Gabrielle Ramus

| Hazard Risk Index | Acceptance Criteria |
| --- | --- |
| 1 to 5 | Unacceptable |
| 6 to 9 | Undesirable |
| 10 to 16 | Acceptable upon completion of quality assurance review |
| 17 to 20 | Acceptable without review |

| Risk Index Table |
| --- |
| Probability of Occurrence | Severity I Catastrophic (Death, Serious Injury) | Severity II Significant (Reversible Serious Injury) | Severity III Marginal (Inconvenience) | Severity IV Negligible |
| Frequent | 1 | 3 | 7 | 13 |
| Probable | 2 | 5 | 9 | 16 |
| Occasional | 4 | 6 | 11 | 18 |
| Remote | 8 | 10 | 14 | 19 |
| Improbable | 12 | 15 | 17 | 20 |

| FMEA Analysis for Universal Facemask Adapter |
| --- |
| Failure | Function or Component | Failure Mode | Effect on System | Possible Hazards | Risk Index | User Detection Means | Applicable Controls |
| Mechanical Failure | Off-label Usage | Improper handling | Possible transmission of infectious particles to the inside of the mask with adapter | Potential for infection | 4 | Placement of adapter and mask with old possibly infected gloves | Disposal of old PPE before placement of new mask containing an adapter |
|  | Adapter Material | Can not withstand pressure and stress | Pressure/stress on face and improper pressure leads to difficulty breathing | Potential for infection | 10 | Difficulty of inspiration and unwanted pressure on the face | Proper labeling of package along with user warnings |
|  | Adapter Design | Adapter adhesive does not properly conform to the inside of the mask or the contours of the face | Gaps in-between the face and the mask | Potential for infection | 6 | Visual detection | Proper labeling of package with applicable face types for small, medium and large adapters. |
| Material Failure | Adapter Material | Material fails to filter particles larger than 5um | Unwanted passage of particulates into respiratory tract | Potential for infection | 8 | Undetectable | Material must have a pore size less than 5um |
|  | Adapter Material | Pores clog | The product is of no use and must be discarded | User will have difficulty breathing | 8 | Breathing becomes difficult | Change the mask every two hours to maintain effectiveness |
|  | Adapter Material | Contamination prior to use | Unwanted passage of particulates into respiratory tract | Potential for infection | 8 | Undetectable | Storage and shipping in sterile package and instructions on how to apply the adapter to masks maintaining sterility |
| Allergic Reaction | Adapter Adhesive | Hives and/or mild irritation | Incompatible for people sensitive to certain materials | User will need to use an adapter without adhesive | 6 | Visible detection | Adhesive will be hypoallergenic, similar to that used on electrode surfaces and adapters will be available without adhesive |
| Contamination | Adapter | Improperly disposed | Improper disposal leads to greater amounts of contamination | Potential for infection | 10 | Visible | Masks containing an adapter should be disposed of in biohazard waste bins |
|  | Adapter | Improper placement/removal | Exposure to possible contaminants | Potential for infection | 12 | Undetectable | The mask containing the adapter will be worn at all times around infected patients and only removed when in a ‘clean’ area |
|  | Packaging | Improperly packaged leading to contamination prior to use | Exposure to possible contaminants before use | Potential for infection | 14 | Package is not properly sealed | Packaging will be inspected prior to shipment, user should also inspect before wearing |